General

Status Released

CE Certified YES

Size

Dimensions Height 49” x Deep 31” x Length 44”

Weight 475 lbs (215kg) Base Unit and pumps ONLY

Recommendations for Lab Space: Front and Back of system need to be accessible Dedicated table needs to hold 600+lbs to account for samples, stock buffers, and side/top/front shelving Space of 3-4 additional feet on each side of system to allow for stock/buffer bottles and shelves to hold- Will be dependent on Accessories purchase

Computer Requirements

Operating System 64 bit Windows 10/ XP

Memory 8 GB RAM (512 MB Video Memory)

Monitor 1280 x 1024 resolution capability

Additional 2 RJ-45 ports (may require additional network card) Minimum 8 USB 2.0 drives; Minimum: 2 SATA Drive Bays and 1 SATA drive; PCI cards: 1 PCI and 1 PCI express 1.0; Integrated Network Adapter (10/100/1000 Mbps)

Power

Rating: 110-220Vac, 50/60Hz, 880 VA max

Full system max output = 10Amps (includes instruments, pumps, computer and UV components)

Environmental Requirements

Operating Temperature: 4-40°C with 0-80% relative humidity, no condensation