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	General

	Status                                                                                                  Released

	CE Certified                                                                                           YES



	Size

	Dimensions                                                                                        Height 49” x  Deep 31” x  Length 44”

	Weight                                                                                                     475 lbs (215kg) Base Unit and pumps ONLY

	Recommendations for Lab Space:     Front and Back of system need to be accessible 
                                                                  Dedicated table needs to hold 600+lbs to account for samples, stock buffers, and side/top/front shelving 
                                                                  Space of 3-4 additional feet on each side of system to allow for stock/buffer bottles and shelves to hold- 
                                                                  Will be dependent on Accessories purchase

	



	Computer Requirements

	Operating System                                                                               64 bit  Windows 10/ XP

	Memory                                                                                                8 GB RAM  (512 MB Video Memory)

	Monitor                                                                                                1280 x 1024 resolution capability               

	Additional                                                                                             2 RJ-45 ports (may require additional network card)
                                                                                                               Minimum 8 USB 2.0 drives; Minimum: 2 SATA Drive 
                                                                                                               Bays and 1 SATA drive; PCI cards: 1 PCI and 1 PCI express 1.0; Integrated Network 
                                                                                                               Adapter (10/100/1000 Mbps)



	Power

	Rating: 110-220Vac, 50/60Hz, 880 VA max

	Full system max output = 10Amps (includes instruments, pumps, computer and UV components)



	Environmental Requirements

	Operating Temperature: 4-40°C with 0-80% relative humidity, no condensation
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